## Slide 1
Individually Adaptable Automatic QT Detector
Yuriy Chesnokov, Dmitry Nerukh, Robert Glen.
Unilever Centre for Molecular Sciences Informatics, University Chemical Laboratory, Cambridge University
Lensfield Road, Cambridge, CB2 1EW, UK
Abstract
A fully automated method is presented for the 7th Computers in Cardiology challenge of QT interval measurement (division 2,3). It is based on
wavelet transforms for QRS detection and T wave localization providing a fast annotation procedure. Variation of the wavelet transform frequency allows us to achieve reasonable quality for T wave end detection despite various morphologies and signal noise.
1. Introduction
The 7th Computers in Cardiology challenge introduced a topic of great importance: fully automatic measurement of QT intervals with acceptable accuracy for clinical evaluations. QT interval prolongation is associated with severe adverse effects. It predisposes the patient to a disorganized heart rhythm called Torsades de Pointes (TdP), which can degenerate into ventricular fibrillation and death. Automated methods with a clinically acceptable quality for QT interval measurement are required for patient monitoring, providing measurements faster than manual interpretation. There are however disagreements between manual and automatic QT interval determination which must to be resolved [1].
2. Method
The continuous (CWT) and fast wavelet transforms (FWT) were used for automatic annotation of the ECG cardio cycle [2]. The annotation method consists of two phases: QRS detection followed by P, T wave location.
2.1. QRS detection
The first stage of the annotation procedure is the detection of the QRS complexes. To amplify the QRS complex and separate low frequency P and T waves and high frequency noise, the CWT transform is applied at 12Hz with an inverse wavelet. The CWT spectrum obtained is further filtered with FWT using an interpolation filter to remove frequency content below 30Hz and the rest of the spectrum is denoised with a hard threshold using a MINIMAX estimate. The reconstructed ECG signal after denoising contains only spikes with nonzero values at the location of the QRS complexes. From this signal, the PQ junction and J point can be located as the boundaries of the spike (Fig. 1).
3. Results
The QT interval was measured in lead II of the PTB Diagnostic ECG database. The database contains various morphologies of the cardio cycle including myocardial infarction patients, cardiomyopathy and healthy controls. Fig. 4 shows results of annotation of the various ECG cycle morphologies from this database.
2.2. P,T waves detection
After QRS complexes are detected the intervals between them are processed for detection of P and T waves. The ECG interval from the J point to the longest duration of the T wave is transformed with CWT on the 3 Hz frequency using an inverse wavelet. Maximum and minimum values on the CWT plot correspond to boundaries and zero crossing to the peak of the T wave (Fig. 2). The frequency of CWT can be adjusted in the range of 1.5-4 Hz along with the time segment duration for better adjustment to various morphologies of the QT interval. This provides for adaptation of the method to a particular shape of the T wave. The same procedure is applied to find a P wave in the PQ interval. The CWT spectrum frequency for P wave detection is varied in the range of 4-9Hz.
4. Conclusions
A fully automated ECG annotation method is presented. It is fast and simple in implementation and the algorithm is applicable to a wide range of ECG cardio cycle morphologies and robust to various types of noise: baseline wander, muscle artefact and power line interference. The consecutive application of CWT and FWT transforms separates noise and P, T waves from the QRS complex. Variation of the frequency for the CWT transform provides adaptation of the method for a particular T wave morphology.
5. References
Moody GB, Koch H, Steinhoff U. The PhysioNet / Computers in Cardiology Challenge 2006: QT Interval Measurement. Computers in Cardiology 2006;33
Addison PS. Wavelet transforms and the ECG: a review. Physiol Meas 2005;26:155-199.
Christov I. et al. Dataset of manually measured QT intervals in electrocardiogram. BioMed Eng Online 2006;5(31).
The representative beat was selected from the annotated signal and the QT interval was measured. The very first preliminary score achieved in the competition was the highest one: 30.35 in division 2. As new entries were submitted by participants, the best score of the challenge decreased to 17. To improve our score the annotation was performed by adjusting the frequency of the CWT transform of the T wave to resemble more closely the published QT intervals [3] measured by expert cardiologists for this challenge. The score achieved was 17.86, very close to the best one, 17.22, in the 2nd division. On the deadline day of the 4th of September we changed to the 3rd division with the same entry and obtained the best preliminary score. The final score is the best in the 3rd division and 2nd place in division 2.
Fig. 1. QRS annotation procedure. Original signal with its CWT spectrum (upper left plane), four high frequency FWT bands of the transformed CWT spectrum with noisy (red) and denoised (black) signal (right half) and reconstructed signal after denoising (lower left plane) with only QRS spikes.
Fig. 2. T wave location procedure. CWT spectrum (right) of the T wave (left) obtained with inverse wavelet on the 3 Hz frequency. Maximum, minimum and zero crossing is marked with red lines on the CWT spectrum. These positions correspond to the boundaries of the T wave and zero crossing to its centre.
Fig. 3. Adaptation of the method to the particular T wave morphology. Annotation of the T wave obtained with CWT transform performed on the 3.5Hz (left) and on the 2.5Hz (right).
Fig. 4. Annotation results of different ECG cycle morphologies and noises from PTB database.
Using a higher frequency for the CWT transform for T wave detection, a shorter QT interval is measured. Fig. 3 shows annotation results for the T wave detected with a CWT transform at 3.5Hz (left) and 2.5Hz (right). The T wave detected on the 2.5Hz CWT spectrum has wider boundaries compared to the 3.5Hz spectrum.